## Slide 1: Interpretable Assessment of ECG Signal Quality Using Expert Noise Annotations from PTB-XL
Reuben N. Addison

## Slide 2
Executive Summary
Project Goal
Develop interpretable machine learning models to classify ECG signal quality
Leverage PTB-XL expert noise annotations as ground truth
Compare logistic regression with deep learning benchmarks
Enable uncertainty-aware clinical decision support
Current Status
Completed:
Logistic Model A (Basic Features)
Logistic Model B (Enhanced Features)
DL1 (Simple MLP Baseline)
Planned:
DL2 (Hybrid CNN Architecture)
DL3 (CNN-Transformer Ensemble)

## Slide 3
Background
Electrocardiography (ECG) is a noninvasive tool for the diagnosis and monitoring of cardiovascular disease, yet its reliability in routine clinical practice is strongly constrained by signal quality (Wagner et al.,2020)
The Problem
Real-world ECGs get degraded by baseline drift, muscle artifact, and electrode contact loss (typically about 23% of the time)
Poor signal quality obscures diagnostic features and makes automated analysis difficult
Manual inspection is slow and doesn't scale in high-volume settings

## Slide 4
Background (Cont.)
Why existing solutions fall short
Fixed SQI thresholds do not generalize
Whole-record quality labels often miss localized noise
Existing models often optimized for diagnosis, not quality.

## Slide 5
Background (Cont.)
Why this approach stands out
  Literature-grounded design
  Built on established ECG signal quality indices rather than task-specific heuristics
  Localized quality assessment
  Window- and lead-level analysis captures transient artifacts missed by whole-record models
  Probabilistic decision-making
  Adaptive thresholds replace fixed cutoffs used in traditional SQI pipelines

## Slide 6
Background (Cont.)
Data
PTB-XL Dataset: 21,837 clinical 12-lead ECGs, 10s at 500 Hz, expert annotations for 4 artifact types (baseline drift, static noise, burst noise, electrode problems), 77% clean / 23% degraded
Train/Val/Test Split: Official PTB-XL stratified folds - Folds 1-8 training (80%), Fold 9 validation (10%), Fold 10 test (10%), threshold optimization on validation set
Approach
Segment ECGs into short, non-overlapping windows at the lead level
Extract hand-crafted, interpretable signal quality indices (SQIs)
Model clean vs. degraded segments using probabilistic methods

## Slide 7
Motivation
Why This Matters
Automated detection flags bad leads instantly making it easier to immediately re-acquire the signal without waiting for manual review
Improves confidence in arrhythmia detection and reduces missed diagnoses from poor signal quality
Interpretable metrics for clinicians

## Slide 8
Methods
Research Question
Can interpretable, literature-based signal quality metrics reliably discriminate clean from degraded ECG segments under realistic clinical noise conditions?
Method
We model ECG signal quality probabilistically using interpretable SQIs and expert-derived labels.
Preprocessing Pipeline: Raw ECG → Median filter baseline removal → 60 Hz notch filter → 0.5-40 Hz bandpass → Per-lead SQI extraction → Feature aggregation (mean/std across leads)

## Slide 9
Feature Engineering: Signal Quality Indices
Category | Features | Description
Spectral | pSQI, basSQI | QRS band power (5-40 Hz) ratio; baseline drift indicator (0-1 Hz)
Statistical | kSQI, sSQI | Normalized kurtosis and skewness of signal amplitude
Morphological | mean, std, rms, ptp | Amplitude statistics for each lead
Hjorth | mobility, complexity | Signal dynamics: frequency content and waveform irregularity
Enhanced* | iSQI, RR-stats, spec_entropy | R-peak detector agreement, HRV features, spectral entropy
* Enhanced features used only in Model B
Feature Selection: Variance thresholding (>0.001) + Mutual information ranking. Model A: 15 features, Model B: 20 features.
Notes: (Feature design grounded in established ECG signal quality literature: Zhao et al., 2018; Rahman et al., 2022)

## Slide 10
Methods: Model Architecture Overview
Model | Type | Input | Key Technique | Status
Model A | Logistic Reg. | 15 Basic SQIs | L1 regularization, class weighting | ✓ Complete
Model B | Logistic Reg. | 20 Enhanced SQIs | L1 regularization, HRV features | ✓ Complete
DL1 | MLP | 15 Basic SQIs | 128→64 hidden, dropout, BCE loss | ✓ Complete
DL2 | Hybrid CNN | Raw waveform | Multi-scale 1D convolutions, residual | Planned
DL3 | CNN-Transformer | Raw waveform | Ensemble + uncertainty, focal loss | Planned
Progression: Interpretable baselines → Feature-based DL → End-to-end deep learning benchmarks
Note: Logistic regression with L1 regularization was adopted after convergence difficulties with the originally planned fully Bayesian logistic regression with horseshoe priors.

## Slide 11
Results: Completed Models
Model | ROC-AUC | AUPRC | Sensitivity | Specificity | ECE
Model A | 0.559 | 0.275 | 0.509 | 0.596 | 0.260
Model B | 0.556 | 0.276 | 0.561 | 0.516 | 0.259
DL1 (MLP) | 0.555 | 0.289 | 0.577 | 0.505 | 0.312
Key Observations
Feature-based models achieve ~55-56% AUC, suggesting SQI features alone have limited discriminative power
Enhanced features (Model B) do not substantially improve over basic features
DL1 shows similar performance to logistic regression on same feature set
ECE values around 0.26-0.31 indicate calibration challenges
Results motivate need for end-to-end deep learning on raw waveforms (DL2, DL3)

## Slide 12
Results: Per-Class Performance Analysis
Model | Clean Prec | Clean Rec | Clean F1 | Deg. Prec | Deg. Rec | Deg. F1
Model A | 0.793 | 0.532 | 0.637 | 0.274 | 0.560 | 0.368
Model B | 0.779 | 0.548 | 0.643 | 0.262 | 0.509 | 0.346
DL1 | 0.792 | 0.271 | 0.403 | 0.251 | 0.775 | 0.380
Interpretation
Class imbalance (77% clean vs 23% degraded) significantly impacts minority class performance
Degraded class F1 scores are substantially lower (0.35-0.38) than clean class (0.40-0.64)
DL1 shows bias toward predicting degraded (high degraded recall, low clean recall)
Motivates use of focal loss and more sophisticated class balancing in DL2/DL3

## Slide 13
Planned: DL2 & DL3 Architectures
DL2 - Hybrid CNN: Input raw 12-lead ECG (12×5000), 4 conv blocks (32→64→128→256 channels), ResNet-style, AdamW optimizer, early stopping. Expected AUC: 0.75–0.80
DL3 - CNN-Transformer Ensemble: Multi-scale CNN stem (parallel kernels k=3,7,15,31), Squeeze-and-Excitation blocks, Transformer encoder (d=128, 8 heads, 4 layers), deep ensemble (5 models) for epistemic uncertainty, focal loss, conformal prediction. Expected AUC: 0.80–0.85

## Slide 14
Evaluation & Validation Strategy
Metrics: ROC-AUC, AUPRC, Sensitivity/Specificity, ECE, Brier Score
Statistical Testing: DeLong test for AUC comparison, Bootstrap CIs, McNemar's test, Calibration curve analysis
Threshold Strategies: Youden's J (maximize sensitivity + specificity), F2-Optimal (weight recall 2×), Cost-Sensitive (FN:FP = 5:1 for clinical safety)
Artifact-Specific Analysis: Performance stratified by baseline drift, muscle artifact, burst noise, and electrode issues
Benchmarking: Performance compared against literature-based deep learning SQA models

## Slide 15
Broader impact & Future directions
Where this goes next
Real-time signal quality monitoring in ECG devices and acquisition systems
Integration into clinical QC workflows and decision-making pipelines
Equity analysis using demographic data to ensure it works fairly across populations

## Slide 16
Summary & Remaining semester work
What has been done ?
Built comprehensive preprocessing and feature extraction pipeline for PTB-XL
Implemented two logistic regression models with interpretable SQI features
Trained DL1 baseline MLP to compare feature-based deep learning
Established evaluation framework with multiple metrics and threshold strategies
Identified limitation: feature-based models plateau at ~56% AUC
What's Next
Implement DL2 hybrid CNN for end-to-end waveform learning
Develop DL3 CNN-Transformer ensemble with uncertainty quantification
Apply focal loss and conformal prediction for calibrated outputs
Perform comprehensive statistical comparisons (DeLong tests)
Analyze per-artifact-type performance breakdown
Document findings in final report